## FormTitan
Header
Fail Analisis
Internal Quality Control (IQC) ID:
Reagent Blank (RB) ID:
Arsenic (As)
RB (ppb):
Sample weight IQC A
Sample weight IQC B
Blank Conc. (ppb)
Cadmium(Cd)
RB (ppb):
Sample weight IQC A
Sample weight IQC B
Blank Conc. (ppb)
Mercury (Hg)
RB (ppb):
Sample weight IQC A
Sample weight IQC B
Blank Conc. (ppb)
Plumbum(Pb)
RB (ppb):
Sample weight IQC A
Sample weight IQC B
Blank Conc. (ppb) | 0.5
Nama dan Tarikh
Nama
Tarikh

## FormGH
Header
Fail Analisis | 110924
Internal Quality Control (IQC) ID: | IQC SCAP 020924
Reagent Blank (RB) ID: | RB GH B 020924
Arsenic (As)
RB (ppb): | 0.04
Sample weight IQC A | 1.504
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 14.406
Cadmium(Cd)
RB (ppb): | 0.01
Sample weight IQC A | 1.504
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 4.109
Mercury (Hg)
RB (ppb): | 0.025
Sample weight IQC A | 1.504
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 8.635
Plumbum(Pb)
RB (ppb): | 0.923
Sample weight IQC A | 1.504
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 92.93
Nama dan Tarikh
Nama | PERMIT MAISARAH
Tarikh | 2024-09-11

## IQC TITAN
Unit Analisis Tradisional
Laporan Internal Quality Control
Fail Analisis: |  |  | 0
Internal Quality Control (IQC) ID: |  |  | 0
Reagent Blank (RB) ID: |  |  | 0
ARSENIC |  |  |  |  | RB (ppb): | 0 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of As in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<6%)
IQC Sample A | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □ |  | #DIV/0!
Within Control Limit □
IQC Sample B | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □
Within Control Limit □
CADMIUM |  |  |  |  | RB (ppb): | 0 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Cd in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<9%)
IQC Sample A | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □ |  | #DIV/0!
Within Control Limit □
IQC Sample B | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □
Within Control Limit □
MERCURY |  |  |  |  | RB (ppb): | 0 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Hg in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<8%)
IQC Sample A | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □ |  | #DIV/0!
Within Control Limit □
IQC Sample B | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □
Within Control Limit □
PLUMBUM |  |  |  |  | RB (ppb): | 0 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Pb in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<5%)
IQC Sample A | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □ |  | #DIV/0!
Within Control Limit □
IQC Sample B | 0 | 0 |  | 0 | #DIV/0! | Within Warning Limit □
Within Control Limit □
Tandatangan Penganalisa & tarikh |  |  |  | 0 |  |  | 0
Tandatangan Penyemak & tarikh

## IQC GH
Unit Analisis Tradisional
Laporan Internal Quality Control
Fail Analisis: |  |  | 110924
Internal Quality Control (IQC) ID: |  |  | IQC SCAP 020924
Reagent Blank (RB) ID: |  |  | RB GH B 020924
ARSENIC |  |  |  |  | RB (ppb): | 0.04 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of As in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<6%)
IQC Sample A | 1.504 | 14.406 | 4563.472 | 4549.066 | 91.22% | Within Warning Limit □ |  | 0.9731%
Within Control Limit □
IQC Sample B | 1.505 | 14.406 | 4519.421 | 4505.015 | 90.4% | Within Warning Limit □
Within Control Limit □
CADMIUM |  |  |  |  | RB (ppb): | 0.01 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Cd in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<9%)
IQC Sample A | 1.504 | 4.109 | 273.695 | 269.586 | 90.1% | Within Warning Limit □ |  | 1.655%
Within Control Limit □
IQC Sample B | 1.505 | 4.109 | 269.271 | 265.162 | 88.68% | Within Warning Limit □
Within Control Limit □
MERCURY |  |  |  |  | RB (ppb): | 0.025 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Hg in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<8%)
IQC Sample A | 1.504 | 8.635 | 454.813 | 446.178 | 89.47% | Within Warning Limit □ |  | 0.4511%
Within Control Limit □
IQC Sample B | 1.505 | 8.635 | 452.805 | 444.17 | 89.13% | Within Warning Limit □
Within Control Limit □
PLUMBUM |  |  |  |  | RB (ppb): | 0.923 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Pb in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<5%)
IQC Sample A | 1.504 | 92.93 | 9639.985 | 9547.055 | 95.73% | Within Warning Limit □ |  | 1.556%
Within Control Limit □
IQC Sample B | 1.505 | 92.93 | 9789.672 | 9696.742 | 97.29% | Within Warning Limit □
Within Control Limit □
Tandatangan Penganalisa & tarikh |  |  |  | PERMIT MAISARAH |  |  | 2024-09-11
Tandatangan Penyemak & tarikh